Prototype of the electromagnetic calorimeter for the Super Charm-Tau factory
Denis Epifanov (BINP, NSU)
Novosibirsk, March 10th 2020
Outline:
- Introduction
- Calorimeters based on CsI(Tl), problems at Super Flavor factories
- Pure CsI endcap calorimeter for Belle II, photopentode/APD options
- Proposal of the calorimeter for Super Charm-Tau factory
- Summary

---
Introduction (I)
Large fraction of π0(→γγ) among the produced hadrons, necessity to reconstruct γ's
in such golden modes as τ→μγ requires a high resolution electromagnetic
calorimeter, which detects γ's in the wide energy range: 10 MeV – 3 GeV
The main tasks for the calorimeter
- High efficiency detection of γ with good energy and coordinate resolutions
- Electron/hadron separation
- Provides signal for the trigger of the detector
- Online/offline luminosity measurement
Full absorption calorimeter based on the fast scintillation crystals with large light yield (LY) is one of the main approaches
Requirements to the calorimeter
- Thick calorimeter to provide good energy resolution in the wide energy range: (16 – 18)X0
- Minimize the passive material in front of the calorimeter: < 0.1X0
- Good time resolution to suppress beam background: < 1 ns
- Fast scintillator (small shaping time) to suppress pileup noise

---
Introduction (II)
- CsI(Tl) has the largest LY, small scintillation decay time and modest price (~3$/cm3). It is used in the electromagnetic calorimeters of modern particle detectors: Belle, Belle II, BaBar, BES-III, CMD-3.
- Lu2SiO5 (LSO), LuAlO3, LYSO are also very good (and much faster than CsI(Tl)), however they are essentially more expensive ((15 – 30)$/cm3), COMET (2000 LYSO crystals).
- Pure CsI has still notable LY, fast decay time component of 30 ns and acceptable price (~5$/cm3). The are several crystal-growing companies which are able to produce needed number of large size crystals (~40 tons): AMCRYS(Ukraine), Saint Gobain (France), HPK (Japan-China), SICCAS (China) → attractive variant for the Super Flavor factories.

---
Belle II electromagnetic calorimeter (ECL)
Belle II ECL is based on 8736 CsI(Tl) crystals (40 tons) with the thickness of 16X0 (30 cm).
It is located inside magnetic field of 1.5 T and covers the solid angle of 91% of 4π.
- Electronics with pipe-line readout and waveform analysis (in the 16-ch Shaper-DSP board) has been developed. It is successfully being exploited now at Belle II.
- To decrease notable pileup noise by a factor of √(1000 ns/30 ns)=5.5 in the endcap ECL (1152+960 ch), CsI(Tl) crystals are planned to be changed to pure CsI crystals.
σE/E ≈ 1.8% (E = 1GeV)
σx = 6 mm/√E(GeV)

---
Belle II endcap ECL upgrade
- To decrease pileup noise by a factor of 5.5 in the endcap ECL, CsI(Tl) are planned to be changed to pure CsI crystals. R&D with CsI(pure) crystals and Hamamatsu photopentodes (PP) showed good results:
  - Low pileup noise, good energy and spatial resolution
  - Similar physical characteristics (as for CsI(Tl)), better radiation hardness
  - There are several crystal producers, acceptable price
- However there are some difficulties: no redundancy, strong dependency on magnetic field, completely new mechanical support is needed. To solve these difficulties second R&D option was suggested: CsI(pure) + Si APD
- In the CsI(pure) + Si APD option we investigated Hamamatsu APD: S8664-1010 and S8664-55.
- With the actual size crystal and 1 APD (1 x 1 cm2) Hamamatsu S8664-1010 we obtained ENE ≈ 2 MeV, while the required ENE ≤ 0.4 MeV
- The main task was to reach admissible level of the electronic noise and the light output of the counter. The wavelength shifter with the nanostructured organosilicon luminophore (NOL-9) is used to improve the light output of the counter by a factor of ~4.
Hamamatsu
APD S8664-55

---
CsI(pure)+WLS+4APD option (I)
- The first tests showed that for the counter, based on the 6 x 6 x 30 cm3 CsI(pure) crystal (AMCRYS) and 1 APD Hamamatsu S8664-1010 (1 cm2, CAPD = 270 pF) coupled to the back facet of the crystal with optical grease (OKEN-6262A) has the light output LO = 26 ph.el./cm2/MeV (for the shaping time of 30 ns), which corresponds to ENE ≈ 2 MeV. Such a small LO and large ENE substantially degrade the energy resolution of the calorimeter (σE/E (100 MeV) ≈ 8%). The acceptable parameters are: LO ≥ 150 ph.el./MeV, ENE < 0.4 MeV → σE/E (100 MeV) = 3.7% (3.4% from the fluctuations of the shower leakage)
- The reason of the small LO: small sensitive area of APD (1/36 of the area of the crystal facet), small quantum efficiency ((20 – 30)%) for the UV scintillation light (320 nm). The reason of large ENE = ENC/LO: small LO and large ENC (large capacitance of Hamamatsu S8664-1010, small shaping time τ = 30 ns →thermal noise ~CAPD/(√τ * gFET) dominates).
- The ways to improve LO and ENE:
  - Increase the number of APDs (LO ~ NAPD, ENE ~ 1/√NAPD) → too expensive
  - Use smaller area APDs: 4 APDs S8664-55 (0.25 cm2, CAPD = 85 pF) (LO is the same, ENE is smaller by a factor of 1/√NAPD = 0.5)
  - Apply wavelength shifter (320 nm → 600 nm)
  - Optimize the input circuit of the preamplifier (increase gFET)
We chose the configuration: CsI(pure) + WLS(nanostructured organosilicon
luminophores) + 4APD (Hamamatsu S8664-55)

---
CsI(pure) + WLS + 4APD option (II)
H. Aihara et al., PoS PhotoDet 2015 (2016) 052.
H. Aihara et al., PoS ICHEP 2016 (2016) 703.
Y. Jin et al., NIMA 824 (2016) 691.
Based on the nanostructured organosilicon luminophores (NOL-9,10,14) from LumInnoTech Co., the WLS plates were developed ((60 x 60 x 5) mm3).
PLQY=95%
τ=7 ns
302 nm → 502 nm
327 nm → 588 nm
337 nm → 655 nm
The absorption and emission spectra of these NOL's match our needs very well (λCsI = 320 nm).
The improvement of the APD QE is by a factor of 2–3.

---
CsI(pure) + WLS + 4APD option (III)
gain(APD) = 50
1300 el
30 ns
Additional
noise
Shot noise
Thermal noise

---
CsI(pure) + WLS + 4APD option (IV)
Crystals, WLS plates and APDs
- We are constructing calorimeter prototype made of 16 counters, the parameters of available crystals (of 6 x 6 x 30 cm3 size) were measured, mechanics was developed, produced and assembled.
- 64 Hamamatsu S8664-55 APDs were purchased from LHC CMS calorimeter group, baking procedure was held at CERN, the dark current was decreased by a factor of about 2.
- 16 WLS plates were purchased, APDs were coupled to the side edges of WLS plates with help of BC-600 optical epoxy resin. The WLS plates with APDs were tested in reference counter.

---
CsI(pure) + WLS + 4APD option (V)
4-channel preamplifier and Shaper-ADC board
- 4-channel charge sensitive preamplifier on 53 x 55 mm2 PCB
- Each channel: sensitivity of 0.2 V/pC, 2 input FET 2SK932 (high transconductance), differential output, HV bias circuit, test pulse input
- 4-channel CAMAC Shaper-ADC board
- CR-(RC)4 filter (τ = 30 ns) + 40 MHz 12-bit pipelined ADC + 256-word circular buffer
- To comply with the new 4-ch preamp additional differential receiver and summator (DRS) boards have been produced and mounted in the Shaper-ADC boards

---
CsI(pure) + WLS + 4APD option (VI)
Development of new electronics for the calorimeter
- Pipeline readout, on-board waveform analysis approach (successfully realized at Belle II ECL)
- Preamplifier is located in the counter, shaping digitization and analysis is implemented in the VME 9U Shaper-DSP board located nearby the detector. Shaper: CR + (RC)4 with the shaping time of 30 ns. Amplitude, time and pedestal are fitted in FPGA of the Shaper-DSP board. The data from the Shaper-DSP boards are sent to the DAQ via optical link (directly or via intermediate collector board)
- The temperature variation of the LY of CsI(pure) is 1.5%/oC, hence, thermostabilization of the calorimeter is needed, the temperature map should be monitored with the accuracy of (0.1 – 0.2) oC

---
CsI(pure) + WLS + 4APD option (VII)
%sigma_{E} over E = {1.9"%"} over nroot{4}{E[GeV]} oplus {{Stat} over {sqrt{E[GeV]}}} oplus { { Elec } over { E[GeV] } }
fluctuation of e/m shower leakage
statistics of
photoelectrons
electronic
noise
F = 1.69 ± 0.04
Stat = 100"%" cdot sqrt{ F over {S[ph.e/MeV] cdot N_{APD} cdot 1000} }
S·NAPD = (160 ± 9) ph.el./MeV
ENE = (0.33 ± 0.03) MeV
Elec = 100"%" cdot { {ENE[MeV]} cdot sqrt{ N_{crys} } over 1000 }
Ncrys = 10 – number of crystals in the cluster

---
Test beam study of 1 counter
- In June 2019 we performed test beam studies with 1 fully assembled CsI counter.
- Electron beam with the energies 0.8, 1.5, 2.0, 2.5 and 3.0 GeV hit the center of the pure CsI crystal. Signals from the counter were recorded with the trigger from the external plastic scintillation detector and from CsI counter itself.
- In total about 1.2M events were recorded, got expected energy deposition spectra from the counter.

---
Super Charm-Tau calorimeter layout
68 counters
in 1 sector
- Crystal of truncated pyramidal form (small facet ~(5.5 x 5.5) cm2) with the length of 30/34 cm (16/18 X0)
- The barrel part includes 5248 counters = 41 θ-rings x 128 counters, total weight is 26/31 tons
- Two endcap parts: 2 x 16 sectors x 68 = 2 x 1088 = 2176 counters, total weight is 10/12 tons
- The whole calorimeter: 7424 counters with the total weight of 36/43 tons → 40/47 M$
- Photodiodes: 7424 → 2 M$
- Electronics: 7424 → 4 M$
- Total price (16X0 / 18X0): 46/53 M$

---
5-month progress in CsI(pure)+WLS+4APD option
- 16 WLS plates with NOL-9 luminophores were purchased, measurement of their characteristics was completed.
- Noise characteristics of 17 4-channel preamplifiers were measured.
- Assembly of 16 counters of the prototype was done, main characteristics were measured. Cosmic runs with the prototype have been started.

---
Summary
- CsI(pure) is an appropriate material for the calorimeter of the Super Charm-Tau factory.
- Beam tests of the prototype based on CsI(pure) crystals and vacuum photopentodes showed good energy and spatial resolutions, as well as essential suppression of the pileup noise.
- The CsI(pure)+WLS+4APDs option is also quite promising. The 16-counter calorimeter prototype has been constructed. Due to the small light yield of the utilized CsI(pure) crystals and ~1.4 larger ENC the LO of the counters are smaller than that of the reference counter (with good CsI(pure) crystal) and ENE is quite high.
- Cosmic runs have been started with the prototype, test beam study of the prototype at ROKK-1M facility in BINP is planned in 2020

---
Backups

---
Study of radiation hardness of CsI(pure) crystals
A. Boyarintsev et al., JINST11 (2016) P03013.
I. Bedny et al., NIMA598 (2009) 273.
- We studied the radiation hardness of 4 CsI(pure) crystals and 1 counter (CsI(pure) + photopentode), they were irradiated by bremsstrahlung γ's with Eγ < 1.4 MeV
- The dose rate was controlled by ELV-6 current and measured by a special dosimeter made of CsI(Tl) crystal and PIN PD
- For the dose of 15 krad the degradation of the LO of 3 crystals and counter was less than 15%, but the degradation of the LO of one counter turned out to be about 60%, it was recovered to about 80% within one year. No change if the Fast/Total-ratio was detected within the accuracy of 3%.
- CsI(pure) crystals were also irradiated by neutrons (up to 1012 1/cm2), we didn't detect any LO degradation within the accuracy of 5%
- The procedure to reject CsI(pure) crystals with poor radiation hardness should be developed